一次性使用取样钳钢丝新增供方

等同性对比报告

Substantially Equivalent Discussions Report of steel wire add new supplier for Single-Use Biopsy Forceps

- 目的 Purpose

根据CR00001653，因单一供方供货风险，兼顾成本优化角度考虑，一次性使用取样钳使用的钢丝新增供方杭州JJJJ有限公司。新供方物料完全按照我司图纸要求，和原供方保持一致，材料牌号和标准保持不变。

According to CR00001653, due to the supply risk associated with sole-source procurement as well as cost optimization objectives, steel wire used in the Single-Use Biopsy Forceps add new supplier Hangzhou Jiawo Spring Co., Ltd. The new supplier 's material is in full compliance with our drawings, consistent with the original supplier, and the material designation and standard remain unchanged.

通过对变更点的评估，变更影响产品的物理性能有：外观、尺寸、操作性能、连接强度。针对这些性能已完成了PQ（PQ报告：MTN-1018283，A/0），现通过等同性对比分析来评价变更后的产品性能依然满足性能指标要求，且在临床上使用是安全有效的。

Through the evaluation on the change point, the physical properties of the product affected by the change include: appearance, size, operational performance, and connection strength. The PQ has been completed for these performances (PQ report: MTN-1018283, A/0). The comparative analysis of equivalence is now performed to evaluate that the performance of the changed product still meets the requirements of performance indicators and is safe and effective for clinical use.

- 参考文件 References

| 文件名Document | 文件编号File No |
| --- | --- |
| 规格性能说明 Product Performance Specification | MTN-0020460 (A/3) MTN-0005526(A/6) MTN-BF-PPS-001(A/8) |
| 抽样标准作业指导书 Sampling standard operation procedure | WCK/PI-0736-001(A/3) |
| 变更申请单 Design Change requests | CR00001653 |
| 取样钳性能验证报告 Performance Qualification Report of Biopsy Forceps | MTN-1018283 (A/0) |
| 一次性使用取样钳设计验证报告 DV Report of Single-Use Biopsy Forceps | MTN-LCNBF-DVR-01 (A/0) MTN-LCNBF-DVR-02(A/0) MTN-0024958 (A/1) MTN-0027841(A/2) |

- 等同性对比分析 SE Analysis

一次性取样钳美国和欧盟市场的性能指标相同，选择任意市场均可，故选择一次性取样钳美国市场进行等同对比分析，性能指标比对详见下表：

The performance indicators of Single-Use Biopsy Forceps are identical in the U.S. and EU markets, so either market can be selected. Thus, the U.S. market for Single-Use Biopsy Forceps is chosen for equivalent comparison and analysis. For details of the performance indicator comparison, see the table below:

性能指标对比 Comparison of performance indicators

| 性能 Performance | 美国 US | 欧盟 EU | 比对结果 |
| --- | --- | --- | --- |
| 外观 Appearance | 产品表面应光滑、无污渍、无毛刺、无裂纹、无划痕，Mark线应清晰 Surface should be smooth, no stain, no burr, no cracks, no scratches and marks should be clear | 产品表面应光滑、无污渍、无毛刺、无裂纹、无划痕，Mark线应清晰 Surface should be smooth, no stain, no burr, no cracks, no scratches and marks should be clear | 相同 Same |
| 尺寸 Dimension | 钳头直径±0.15mm Diameter of jaws±0.15mm 工作长度L±25mm Working length L±25mm | 钳头直径±0.15mm Diameter of jaws±0.15mm 工作长度L±25mm Working length L±25mm | 相同 Same |
| 操作性能 Operating performance | 将弹簧管盘圈（直径D约为200mm），然后前端勾头。在匀速运行时，当钳头打开和关闭时，没有明显的阻力和断裂。（注：工作长度≤1600mm盘1圈并勾头；>1600 mm，盘2圈并勾头。钩式取样钳不勾头） Turning the spring sheath for loops (D is about 200mm) and then hook the top. To operate in the uniform speed, there is no obvious resistance and break when jaws are opened and closed. (Note: Working length ≤1600 mm turning 1 loop and hook the top; >1600 mm, turning 2 loops and hook the top. Twin Wire Biopsy Forceps don’t hook the top) | 将弹簧管盘圈（直径D约为200mm），然后前端勾头。在匀速运行时，当钳头打开和关闭时，没有明显的阻力和断裂。（注：工作长度≤1600mm盘1圈并勾头；>1600 mm，盘2圈并勾头。钩式取样钳不勾头） Turning the spring sheath for loops (D is about 200mm) and then hook the top. To operate in the uniform speed, there is no obvious resistance and break when jaws are opened and closed. (Note: Working length ≤1600 mm turning 1 loop and hook the top; >1600 mm, turning 2 loops and hook the top. Twin Wire Biopsy Forceps don’t hook the top) | 相同 Same |
|  | 钳头应能平稳地打开和关闭，无卡顿和断裂现象。 The jaws should be opened and closed smoothly, no jam and break. | 钳头的打开和关闭应平稳，无卡顿和断裂现象。 The jaws should be opened and closed smoothly, no jam and break. | 相同 Same |
|  | 滑块运动平稳，无异常响声，无空行程。 The finger-ring can move smoothly, no abnormal sound, no empty journey in the motion. | 滑块运动平稳，无异常响声，无空行程。 The finger-ring can move smoothly, no abnormal sound, no empty journey in the motion. | 相同 Same |
| 连接强度 Connection strength | 整钳最小拉力为10kgf。 The minimum tension for the finished forceps is 10 kgf. | 整钳最小拉力为10kgf。 The minimum tension for the final forceps is 10 kgf | 相同 Same |

3.1针对变更前、变更后的产品的外观、尺寸和操作性能，通过目视的方式进行确认，按照统一的测试方法及标准，变更后的产品仍然满足技术要求。

The appearance, dimension and operating performance of the product before and after the change are all confirmed visually, in accordance with unified test methods and standards, so the changed product still meets the technical requirements.

| 性能 Performance | 变更前 Before change | 变更后 After product | 结论 Conclusion |
| --- | --- | --- | --- |
| 外观 Appearance | Pass | Pass | 无差异 No difference |
| 尺寸 Dimension | Pass | Pass | 无差异 No difference |
| 操作性能 Operating performance | Pass | Pass | 无差异 No difference |

  - 将产品的连接强度做为两组样本，来推断它们是否存在差异。利用Minitab软件

中基本统计计算P值功能，将变更前变更后产品的数据分别代入计算，求得P值。当P值＜0.05，那么判断两组数据有明显差异，当P值＞0.05，则判断两组数据间无明显差异。

Take the connection strength of the product as two groups of samples infer whether there is a difference between them. Use the basic statistical calculation function of the Minitab software to calculate the P value, and substitute the data of the product before and after the change into the calculation to obtain the P value. When the P value is less than 0.05, it is judged that there is a significant difference between the two groups of data, and when the P value is greater than 0.05, it is judged that there is no significant difference between the two groups of data.

3.2.1连接强度对比分析Comparative analysis of connection strength

| No. | 产品变更前连接强度数据A(kgf) Connection strength data before product A(kgf) | 产品变更后连接强度数据B(kgf) Connection strength data after product change B (kgf) |
| --- | --- | --- |
| 1 | 16.94 | 24.67 |
| 2 | 16.51 | 22.59 |
| 3 | 16.32 | 24.14 |
| 4 | 16.84 | 25.46 |
| 5 | 16.15 | 26.27 |
| 6 | 16.34 | 26.07 |
| 7 | 16.17 | 24.29 |
| 8 | 16.95 | 25.64 |
| 9 | 16.67 | 23.55 |
| 10 | 16.74 | 22.34 |
| 11 | 18.39 | 23.44 |
| 12 | 17.89 | 23.08 |
| 13 | 18.42 | 21.93 |
| 14 | 17.92 | 24.99 |
| 15 | 18.43 | 21.99 |
| 16 | 18.15 | 23.50 |
| 17 | 17.61 | 23.99 |
| 18 | 17.02 | 22.48 |
| 19 | 17.76 | 24.93 |
| 20 | 17.75 | 22.58 |
| 21 | 16.28 | 22.05 |
| 22 | 16.27 | 22.11 |
| 23 | 17.99 | 22.56 |
| 24 | 15.98 | 23.69 |
| 25 | 14.75 | 25.40 |
| 26 | 16.78 | 28.69 |
| 27 | 18.10 | 23.29 |
| 28 | 16.60 | 22.84 |
| 29 | 16.45 | 23.41 |
| 30 | 16.21 | 27.55 |
| 最小值Min | 14.75 | 21.93 |
| 最大值Max | 18.43 | 28.69 |
| 平均值AVG | 17.01 | 23.98 |
| 标准差STD | 0.904 | 1.701 |

备注：变更前的连接强度数据A详见附件1: 一次性使用大口径取样钳设计验证报告MTN-0024958 (A/1），变更后的连接强度数据B详见附件2: 取样钳性能验证报告MTN-1018283 (A/0)。

Remarks: The connection strength data A1 before change is shown in Annex 1: Design Verification Report of Jumbo Single-Use Biopsy Forceps MTN-0024958 (A/1）, connection strength Data B after the change is shown in appendix 2: Performance Qualification Report of Biopsy Forceps MTN-1018283 (A/0).

实验结果分析 Analysis of results:

| DL | 变更前后连接强度数据比对P值 Comparison of connection strength data before and after change | P值判断 P value judgment |
| --- | --- | --- |
| A、B | 0 | 有明显的差异Significant difference |

数据A、B进行分析，具体分析如下

Data A and B are analyzed, the specific analysis is as follows：

通过以上数据分析，变更前连接强度平均值17.01，最小值14.75，最大值18.43，标准差0.904；变更后连接强度平均值23.98，最小值21.93，最大值28.69，标准差1.701。变更前后的连接强度比对P值为0＜0.05，有明显的差异。

Through the above data analysis, the average connection strength before the change is 17.01, the minimum is 14.75, the maximum is 18.43, and the standard deviation is 0.904; After the change, the average connection strength is 23.98, the minimum is 21.93, the maximum is 28.69, and the standard deviation is 1.701. The P value of the connection strength before and after the change is 0<0.05, which is a significant difference.

以上取样钳A、B测试结果判定有差异，根据取样钳整钳连接强度≥10kgf，变更前后连接强度都在规定范围内，所以仍然满足要求。

There is a difference in the test results of Biopsy Forceps, A and B. According to the Biopsy Forceps connection strength is ≥ 10kgf, the connection strength before and after the change is within the specified range, so it still meets the requirements.

# 4. 结论 Conclusion

通过分析证明变更后一次性使用取样钳仍能够满足设计要求且能保证产品的安全有效性。

The analysis proves that the Single-use Biopsy Forceps can still meet the design requirements and ensure the safety and effectiveness of the product after the change.

版本记录Version of the record

| 序号 No | 版本 Version | 日期 Date | 描述 Description | 作者 Author |
| --- | --- | --- | --- | --- |
| 1 | A/0 | 2026-01-08 | 首次编制 First draft | XXQ |